## Slide 1: H2GCROC
20250611

## Slide 2: Hardware
Clock & Trigger & Power cable xN
KCU105 & protoboard x4
Common clock board
Trigger fan out board
High Voltage Supply
Signal generator x1
Crystal
Adapter board

## Slide 3: FAQ
Must insert jumper J14 on the FPGA board
If you have multiple protoboard and FPGA board, you need a common clock input

## Slide 4: Programming FPGA
Install Vivado 2024.2 (no need license key)
Launch Vivado 2024
  Create a project with KCU105 board
  Connect JTAG to PC and power on FPGA board
Click Open target on the left Panel
  Open the device
  Right click on it and Click “Add Configuration Memory Device”
  Select device “mt25qu256-spi-x1_x2_x4” and continue
  Select file “frame_protoboard_v2.mcs” (version v4.11)
  Program it
If successful, the display will be as shown on the right picture
Firmware: v4.11
  https://web.tresorit.com/l/86zmA#BE8fqp4of3N8pFNz-cXpgg&ZZ8bIfjDltJSr5qvdDoZYuGsgz6LDm0Q
Programming guide: KCU_Flashing.pdf
  https://web.tresorit.com/l/86zmA#BE8fqp4of3N8pFNz-cXpgg&GYtZ9jQtioHk1QwqwK5nizBt2rLRRpSt
1
2

## Slide 5: FPGA IP address
DIP Switch on KCU105
FPGA IP address = 10.1.2.208 + DIP SWITCH[2:0]
DIP SWITCH[2:0]
210

## Slide 6: Software H2GConfig
Install a Linux PC with OS AlmaLinux 9.5
Install AnaConda and create a virtual environ
  conda create -n hgc
  conda activate hgc
  conda install -c conda-forge pygame pyside6 numpy matplotlib loguru
H2GConfig
  conda activate hgc
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GConfig.git
  cd ~/workspace/H2GConfig
  python H2GConfig.py

## Slide 7: Software H2GDAQ
The same virtual environ work for H2GDAQ
  conda activate hgc
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GDAQ.git
  cd ~/workspace/H2GDAQ
  git checkout dev-0v12
  python H2GDAQ.py

## Slide 8: Software H2GCalib
Create a virtual environ for H2GCalib or add package to hgc environ
  conda activate hgc
  conda install -c conda-forge pygame pyside6 numpy matplotlib loguru tqdm pandas
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GCalib_3B
  cd ~/workspace/H2GCalib_3B
Launch it with GUI
  python 100_UI.py
Launch it with console
  Python 000_SocketPool.py
  ...
  Python 010_ToTCalib.py

## Slide 9: Ethernet communication
Connect ethernet cable between PC & FPGA board
Set PC IP address to 10.1.2.207
  $ sudo ifconfig enp4s2 10.1.2.207 netmask 255.255.255.0
  $ sudo ifconfig enp4s2 mtu 9000 up
make sure your switch, cable and network card support 1 Gbps speed, 100 Mbps will not work
  $ ethtool enp4s2 | grep Speed
  Speed: 1000Mb/s
H2GConfig has a button “Ping FPGA” can test the connection, it don’t support under windows
Ping
  On Linux: ping 10.1.2.208 -s 32
  On Windows: ping 10.1.2.208 -l 32

## Slide 10: Ethernet communication
Network Optimization
Edit /etc/sysctl.conf and add the following lines:
  net.core.rmem_max=8388608
  net.core.wmem_max=8388608
  net.core.rmem_default=8388608
  net.core.wmem_default=8388608
net.core.netdev_max_backlog=250000
  net.ipv4.udp_mem=8388608 8388608 8388608
  net.ipv4.udp_rmem_min=16384
  net.ipv4.udp_wmem_min=16384
  Then run the following command to apply the changes:
sudo sysctl -p

## Slide 11: Ethernet Firewall
AlmaLinux 9.5
DAQ and Calibration programs use UDP port 11000 ~
[root@localhost ~]# firewall-cmd --zone=public --add-port=11000-11100/udp --permanent
success
[root@localhost ~]# firewall-cmd --zone=public --list-all
public (active)
target: default
icmp-block-inversion: no
interfaces: eno1
sources:
services: cockpit dhcpv6-client ssh
ports: 11000-11100/udp
protocols:
forward: yes
masquerade: no
forward-ports:
source-ports:
icmp-blocks:
rich rules:

## Slide 12: Architectural
72 channels of analog signals input
4 CM (only connect to ADC internally)
2 Calibration channel
320MHz Clock input
Command link (L1A is trigger input)
Two lane for data output:
  Speed 1.28Gbps
  Charge info, Time info…
Four lane for trigger generation
  1.28Gbps x4
I2C interface read/write 7900 parameters
CC: Current conveyor (VFE), calibration SiPM gain variation
PA: Preamplifier
TOT: Time Over Threshold, dedicated discriminator and TDC (50ps), 12 bit
TOA: Time Of Arrival, dedicated discriminator and TDC (24ps), 12 bit
SH: Shaper
ADC: 10bit 40MHz SARADC

## Slide 13: Default register value
The default config value of H2GCROC after power up
"Channel_0-71 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"CM_0-3 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"CALIB_0-1 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"Global_Analog_0-1 ": "cf d3 83 28 28 00 9a 9a a2 8a 50 4b 4b 68 6f",
"Reference_Voltage_0-1 ": "dc 00 3c 1c 58 d0 00 00 00 00 00 00 00",
"Master_TDC_0-1 ": "37 d4 40 80 08 b0 02 00 00 80 08 b0 02 00 00 00",
"Digital_Half_0-1 ": "00 00 00 00 80 00 00 00 00 00 00 00 00 19 00 08 cc cc cc 0c cc cc cc cc 0f 01 00",
"Top ": "08 0f 40 31 00 03 85 00 00 00 00 00 00 00 00 00 00 00 00 00 00"
"HalfWise_0 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
Total 1333 Bytes

## Slide 14: Operation Procedure
H2GCROC can generate a pulse signal internally and inject it to input channel, with adjustable amplitude
H2GCalib_3B
  Set IODELAY – adjust the data/trigger line delay between FPGA and H2GCROC, this parameter is provided for FPGA, not for H2GCROC.
  ADC pedestal
  TOA calibration – adjust the channel threshold to make the output value close to the user's target value
  TOT calibration – adjust the channel threshold to make the output value close to the user's target value
Repeat this procedure for all FPGA and H2GCROC combination and generate configuration files

## Slide 15: Operation Procedure
H2GDAQ
  Set IODELAY – this step is the same as calibration
H2GConfig
  Send All configs – upload all calibration values to all H2GCROC
H2GDAQ
  Start DAQ
Decode data or Online Monitoring

## Slide 16: H2GCalib_3B
Modify file H2GCalib_3B/config/common_settings_3B.json
{
"udp": {
"h2gcroc_ip": "10.1.2.208",
"pc_ip": "10.1.2.207",
"h2gcroc_port": 11000,
"pc_cmd_port": 11000,
"pc_data_port": 11001
}
}

## Slide 17: H2GCalib_3B – Socket Pool
python H2GCalib_3B/ 000_SocketPool.py
This program maintains multiple network connections, and other calibration programs use this as a relay, keep this program running in the background
[Pool] Control → 127.0.0.1:6002
[Pool] Data → 127.0.0.1:6001
...

## Slide 18: H2GCalib_3B – IO delay
python H2GCalib_3B/007_IODelay.py
You can also use H2GDAQ’s IO Delay button to do this, but it don’t generate report

## Slide 19: H2GCalib_3B – ADC pedestal calibration
This program try to get the best pedestal value, it start from config “default_2024Aug_config.json”
python H2GCalib_3B/005_PedestalCalib.py
Before calibration ADC pedestal
After calibration ADC pedestal
avg
target

## Slide 20: H2GCalib_3B – ToA Calibration
python H2GCalib_3B/006_ToACalib.py -t 50 -i ./dump/005_PedestalCalib_data_20250618_183209/config_pede_a0_20250618_183343.json,./dump/005_PedestalCalib_data_20250618_183209/config_pede_a1_20250618_183343.json
Before calibration ToA
After calibration ToA

## Slide 21: H2GCalib_3B – ToT Calibration
python H2GCalib_3B/010_ToTCalib.py -t 300 -i ./dump/006_ToACalib_data_20250618_205115/config_toa_a0_20250618_212606.json,./dump/006_ToACalib_data_20250618_205115/config_toa_a1_20250618_212606.json
Before calibration ToT
After calibration ToT
channel broken
underflow

## Slide 22: H2GCalib_3B – inject 2V5 signal
python.exe .\009_IntInjection_2V5.py -i .\dump\010_ToTCalib_data_20250618_213759\config_tot_a0_20250618_221027.json,.\dump\010_ToTCalib_data_20250618_213759\config_tot_a1_20250618_221027.json
ASIC 0, half 0 is broken (ch 1 ~ 9)
ASIC 1 is fine (ch 77 ~ 85)

## Slide 23: Operation Procedure - DAQ
Directory to store data
DAQ stop condition: By event count, by time or manual
Start/Stop DAQ button
high voltage relay setting
A0 = ASIC 0 only
A1 = ASIC1 only
F0 = ASIC0 and ASIC1 both
Send command to turn On/Off the relay base on the settings above. You also need to turn on your power supply
1
6
5

## Slide 24: Operation Procedure - Config
H2GConfig
  The GUI can read the configuration file and display it graphically. It allow you to modify register bit by bit
  It can upload configuration to the ASIC
  There is also a Ping FPGA button for you to test the ethernet connection
Test network connection
Load configuration
Upload to ASIC
3
2
4

## Slide 25: Operation Procedure - online_monitoring
https://github.com/tlprotzman/h2g_online_monitoring
dependence: ROOT
developed by Tristan Protzman from Lehigh University
$ git clone https://github.com/tlprotzman/h2g_online_monitoring.git
$ cd ~/workspace/h2g_online_monitoring
$ make
$ export DATA_DIRECTORY=~/workspace/H2GDAQ
$ export MONITORING_PORT=12345
$ root -q -b -x -l RunMonitoring.c\(RunNumber\)
Processing RunMonitoring.c(9)...
Parsing config file eeemcal_20sample.cfg
Key: MAX_SAMPLES, Value: 20
Key: NUM_FPGA, Value: 4
Key: DETECTOR_ID, Value: 2
Configuring for detector ID 2: EEEMCal
Configuration values:
NUM_LINES: 5
NUM_FPGA: 4
NUM_CHANNELS: 72
NUM_LINES: 5
MAX_SAMPLES: 20
MACHINE_GUN_MAX_TIME: 600
PACKET_SIZE: 1452
Configuring for detector ID 2: EEEMCal
Configuration values:
NUM_LINES: 5
NUM_FPGA: 4
NUM_CHANNELS: 72
NUM_LINES: 5
MAX_SAMPLES: 20
MACHINE_GUN_MAX_TIME: 600
PACKET_SIZE: 1452
huh??
Info in <THttpEngine::Create>: Starting HTTP server on port 12345
Attempting to open file /home/shyao/workspace/H2GDAQ/Run009.h2g
Starting at byte 828
Building events... done!
Updating canvases... done!
.....
$ cat eeemcal_20sample.cfg
MAX_SAMPLES=20
NUM_FPGA=4
DETECTOR_ID=2

## Slide 26: Operation Procedure - online_monitoring
Launch web browser and connect to localhost:12345
before calibration
After calibration

## Slide 27: Operation Procedure - online_monitoring
Inject light to crystal, each crystal use 4x4 SiPM. The light comes from an LED with a pulse width of 100ns
Light inject to this 16 channels
Light leaking to other crystal, must improve isolation
Crosstalk and Light leaking

## Slide 28: Operation Procedure - online_monitoring
Draw in crystal Array, it depend on your channel mapping
Light inject to this crystal
Light leaking to other crystal, must improve isolation
Other channel is not connected during test

## Slide 29: Operation Procedure - online_monitoring
Inject light to crystal, without leaking
Light inject to this 16 channels
Crosstalk (in H2GCROC)

## Slide 30: Operation Procedure - decode
https://github.com/tlprotzman/h2g_decode
dependence: ROOT
developped by Tristan Protzman from Lehigh University
It provide a library h2g_decode and a executable h2g_run
$ cd workspace
$ git clone https://github.com/tlprotzman/h2g_decode.git
$ cd ~/workspace/h2g_decode
$ mkdir data
$ mkdir build; cd build; cmake ..; make
$
$ export DATA_DIRECTORY= ~/workspace/H2GDAQ-main
$ export OUTPUT_DIRECTORY=~/workspace/h2g_decode
$ ./h2g_run
** Data file is named with “Run%03d.h2g”

## Slide 31: Operation Procedure - decode
To install root under Linux and anaconda virtual environment
$ conda create -n myenv
$ conda activate myenv
$ conda install -c conda-forge root
$ cd workspace/
$

## Slide 32: Q/A
WWW

## Slide 33: Other
Use EEPROM/MRAM to store calibration value? (radiation harden)