## Slide 1: Weekly meeting
Yu-Siang Xiao (蕭宇翔)
2026 06/30

## Slide 2: Trigger Matching Between H2GCROC and CITIROC
Goal: Match events recorded by two DAQ systems.
Method based on inter-event time intervals.
Different DAQ systems have different event numbering.
Timestamp comparison is more robust than raw event ID.

## Slide 3: Experiment setup
Det1: LYSO 8.0X0
SiPM(Osemi 60035 @ 29.0V)
H2GCROC
  External trigger by CITIROC
Det2: GAGG 0.5X0
SiPM(Hamamatsu @ HV: 29|10 ~ 56V)
GTM-EIC-CITIROC
  Self-trigger
Source: 22Na:
2 x 511 keV(e+-e- back to back γ)
1274 keV γ
Det1
H2GCROC
LYSO 8X0
Det2
CITIROC
GAGG 0.5X0
Trigger
signal
LYSO
GAGG
511 keV γ
511 keV γ
1274 keV γ
May get signal on each detector in same time, to try the event matching!
Trigger
signal

## Slide 4: Procedure
5,6
To make sure 2 detector take the same events. ex: 30 sec
1. Set the high THR(DEC=1k) of trigger detector:
  Cut all of the data, make the trigger index reset.
2. Open the data taking of H2GCROC and CITIROC.
  Now, the data flow, and trigger signal is clear!
3. Set the lower THR of trigger detector:
  Start to detector signal and sent the trigger.
  The CITIROC(self trigger) and H2GCROC(external trigger) start to take the data save.
4. Wait for 30 sec!
5. Set the high THR of trigger detector:
  Close the trigger and stop the data flow.
6. Close and save the data flow.
The event # of CITIROC should equal to trigger # of H2GCROC.
LYSO
GAGG
LYSO
GAGG
Trigger
signal
LYSO
GAGG
Trigger
signal
LYSO
GAGG
Trigger
signal
1
2
3,4

## Slide 5: Time Interval Extraction
Because CITIROC sent trigger to H2GCROC, # Event of CitiROC>H2GCROC
CITIROC: , (use fine time counter “fcnt”)
H2GCROC: , (use “timestamp_r”)
Build vectors by Event sequence for 2 rocs data:
  Vector(dtC) of CITIROC
  Vector(dtH) of H2GCROC
tH0
tH1
tH3
tH4
tH5
dtH0
dtH1
dtH2
dtH3
tC0
tC1
tC2
tC3
tC4
tC5
dtC0
dtC1
dtC2
dtC3
dtC4

## Slide 6: Best time window Search
Estimate corresponding index by proportional position.
Search only ±1% window.
Choose an appropriate minimum |dtH-dtC|.
Scan: dtC select ith , dtH select jth:
  PS: ith at R =
  Total time difference: D =
Fill D into the histogram:
  The real physic event will located in a range ~ 0ns.
  Not matched event will spread in random difference.
ddt~0=real event, matched i,j
ddt=reandom number
Not matched i,j
dtH0
dtH1
dtH2
dtH3
dtC0
dtC1
dtC2
dtC3
dtC4
dtHm
dtCn
…
…

## Slide 7: Matching Conditions
Take the matched event: |Δt difference| < 240 ns.
Prevent duplicate matches using used_H / used_C.
Store both previous and current events to be matched event.
Zoom in

## Slide 8: ADC Extraction
Reload matched CITIROC / H2GCROC events.
Fill ADC(peak) in both system.
Fill SumADC in H2GCROC().
H2GCROC
511keV
1274keV
511keV
1274keV
511keV
1274keV

## Slide 9: QA for the spectrum
Why the second peak(1274keV) has a large difference statistics?
  The total length between LYSO(8X0) and GAGG(0.5X0), Cause different interaction rate of 1274 keV
Why does the H2CROC system has a lower efficiency?
  The solid angle between detectors. And the ROCs data matching problem will loss the data.
LYSO
GAGG
511 keV γ
QA for the matching algorithm
How much background we label them to be matched?
background
background
signal
signal
Zoom in
and select
+-240 ns
+-240ns
+-240 is 48 bins
Signal ~ 48*(1700-800)/2
BG ~ 48*800
S/N = 450/800=56%
Real event ratio: 450/1250=36%
Need to improve the algorism?
Or maybe the 2 trigger system could reject BG?