## Slide 1: H2GCROC
2025 - 2026

## Slide 2: Hardware
Clock & Trigger & Power cable xN
KCU105 & protoboard x4
Common clock board
Trigger fan out board
High Voltage Supply
Signal generator x1
Crystal
Adapter board

## Slide 3: FAQ
Must insert jumper J14 on the FPGA board
If you have multiple protoboard and FPGA board, you need a common clock input

## Slide 4: Programming FPGA
Install & Launch Vivado (no need license key)
Open Hardware Manager
Connect JTAG to PC and power on FPGA board
  Click Open target on the left Panel
  Open the device
  Right click on it and Click “Add Configuration Memory Device”
  Select device “mt25qu256-spi-x1_x2_x4” and continue
  Select file “frame_protoboard_v2.mcs” (version v4.11)
  Program it
Firmware: v4.11
  https://web.tresorit.com/l/86zmA#BE8fqp4of3N8pFNz-cXpgg&ZZ8bIfjDltJSr5qvdDoZYuGsgz6LDm0Q
Programming guide: KCU_Flashing.pdf
  https://web.tresorit.com/l/86zmA#BE8fqp4of3N8pFNz-cXpgg&GYtZ9jQtioHk1QwqwK5nizBt2rLRRpSt

## Slide 5: FPGA IP address
DIP Switch on KCU105
FPGA IP address = 10.1.2.208 + DIP SWITCH[2:0]
DIP SWITCH[2:0]
210

## Slide 6: Software environ
AlmaLinux 9.5, Use AnaConda
  conda create -n hgc
  conda activate hgc
  conda install -c conda-forge pygame pyside6 numpy matplotlib loguru pandas tqdm
Ubuntu 24.04, Use python venv
  sudo apt-get install python3-pip python3-venv libxcb-cursor-dev net-tools
  mkdir myenv
  python3 -m venv myenv
  source myenv/bin/activate
  pip install pyside6 numpy matplotlib loguru pygame pandas tqdm
Windows 11, Use AnaConda
  the same as AlmaLinux 9.5, Use AnaConda, you can install git with conda
baremetal, other ...

## Slide 7: Software H2GConfig
H2GConfig
  conda activate hgc
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GConfig.git
  cd ~/workspace/H2GConfig
  python H2GConfig.py
5e478442b 2025/2/26

## Slide 8: Software H2GDAQ
The same virtual environ work for H2GDAQ
  conda activate hgc
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GDAQ.git
  cd ~/workspace/H2GDAQ
  git checkout dev-0v12
  python H2GDAQ.py
8a17963 2025/4/25

## Slide 9: Software H2GCalib
Create a virtual environ for H2GCalib or add package to hgc environ
  conda activate hgc
  conda install -c conda-forge pygame pyside6 numpy matplotlib loguru tqdm pandas
  cd workspace
  git clone https://gitlab.cern.ch/sjia/H2GCalib_3B
  cd ~/workspace/H2GCalib_3B
Usage 1: GUI mode
  python 100_UI.py
Usage 2: console mode
  Python 000_SocketPool.py
  ...
  Python 010_ToTCalib.py
0e179be2b 2025/10/12
8dce2723fb 2025/5/9

## Slide 10: Ethernet communication
Connect ethernet cable between PC & FPGA board
Set PC IP address to 10.1.2.207
  $ IF=enp4s2
  $ sudo ifconfig ${IF} 10.1.2.207 netmask 255.255.255.0
  $ sudo ifconfig ${IF} mtu 9000 up
make sure your switch, cable and network card support 1 Gbps speed, 100 Mbps will not work
  $ sudo ethtool ${IF} | grep Speed
  Speed: 1000Mb/s
H2GConfig has a button “Ping FPGA” can test the connection, it don’t support under windows
Ping
  On Linux: ping 10.1.2.208 -s 32
  On Windows: ping 10.1.2.208 -l 32

## Slide 11: Ethernet communication
Network Optimization (optional)
Create config file 50-h2gcroc.conf and add the following lines:
  net.core.rmem_max=8388608
  net.core.wmem_max=8388608
  net.core.rmem_default=8388608
  net.core.wmem_default=8388608
  net.core.netdev_max_backlog=250000
  net.ipv4.udp_mem=8388608 8388608 8388608
  net.ipv4.udp_rmem_min=16384
  net.ipv4.udp_wmem_min=16384
Then run the following command to apply the changes:
  sudo cp 50-h2gcroc.conf /etc/sysctl.d/
  sudo sysctl -p

## Slide 12: Ethernet firewall 2
Almalinux 9.5
Use external and internal firewall zone
$ firewall-cmd --zone=public --remove-interface=enp87s0
$ firewall-cmd --zone=internal --add-interface=enp87s0
$ firewall-cmd --zone=internal --add-port=11000-11100/udp
$ firewall-cmd --runtime-to-permanent
$ firewall-cmd --zone=internal --list-all
internal (active)
target: default
icmp-block-inversion: no
interfaces: enp87s0
sources:
services: cockpit dhcpv6-client mdns mountd nfs nfs3 rpc-bind samba samba-client ssh
ports: 11000-11100/udp
protocols:
forward: yes
masquerade: no
forward-ports:
source-ports:
icmp-blocks:
rich rules:
Note interface enp86s0 connect to office network interface enp87s0 is for h2gc data transfer
Disable SELinux

## Slide 13: multiple IP
Setting PC IP address for 10G device
For testing, assign multiple IP addresses to a single interface
$ nmcli dev status
$ nmcli con down "enp87s0"
$ nmcli con modify "enp87s0" ipv4.addresses "10.1.2.207/24" ethernet.mtu 9000 ipv4.method manual
$ nmcli con up "enp87s0"
AlmaLinux 9.5
$ nmcli con down "enp87s0"
$ nmcli con modify "enp87s0" ipv4.addresses "10.1.2.207/24,10.1.2.206/24,10.1.2.205/24,10.1.2.204/24" ethernet.mtu 9000 ipv4.method manual
$ nmcli con up "enp87s0"
$ nmcli dev status
$ nmcli con show

## Slide 14: Architectural
72 channels of analog signals input
4 CM (only connect to ADC internally)
2 Calibration channel
320MHz Clock input
Command link (L1A is trigger input)
Two lane for data output:
  Speed 1.28Gbps
  Charge info, Time info…
Four lane for trigger generation
  1.28Gbps x4
I2C interface read/write 7900 parameters
CC: Current conveyor (VFE), calibration SiPM gain variation
PA: Preamplifier
TOT: Time Over Threshold, dedicated discriminator and TDC (50ps), 12 bit
TOA: Time Of Arrival, dedicated discriminator and TDC (24ps), 12 bit
SH: Shaper
ADC: 10bit 40MHz SARADC

## Slide 15: Default register value
The default config value of H2GCROC after power up
"Channel_0-71 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"CM_0-3 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"CALIB_0-1 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"Global_Analog_0-1 ": "cf d3 83 28 28 00 9a 9a a2 8a 50 4b 4b 68 6f",
"Reference_Voltage_0-1 ": "dc 00 3c 1c 58 d0 00 00 00 00 00 00 00",
"Master_TDC_0-1 ": "37 d4 40 80 08 b0 02 00 00 80 08 b0 02 00 00 00",
"Digital_Half_0-1 ": "00 00 00 00 80 00 00 00 00 00 00 00 00 19 00 08 cc cc cc 0c cc cc cc cc 0f 01 00",
"HalfWise_0 ": "00 00 00 00 00 00 00 00 00 00 00 00 00 00 00",
"Top ": "08 0f 40 31 00 03 85 00 00 00 00 00 00 00 00 00 00 00 00 00 00"
Total 1363 Bytes

## Slide 16: Operation Procedure
H2GCROC can generate a pulse signal internally and inject it to input channel, with adjustable amplitude
H2GCalib_3B
  Set IODELAY – adjust the data/trigger line delay between FPGA and H2GCROC, this parameter is provided for FPGA, not for H2GCROC.
  ADC pedestal
  TOA calibration – adjust the channel threshold to make the output value close to the user's target value
  TOT calibration – adjust the channel threshold to make the output value close to the user's target value
Repeat this procedure for all FPGA and H2GCROC combination and generate configuration files

## Slide 17: Operation Procedure
H2GDAQ
  Set IODELAY – this step is the same as calibration
H2GConfig
  Send All configs – upload all calibration values to all H2GCROC
H2GDAQ
  Start DAQ
Decode data or Online Monitoring

## Slide 18: H2GCalib_3B
Modify file H2GCalib_3B/config/common_settings_3B.json
{
"udp": {
"h2gcroc_ip": "10.1.2.208",
"pc_ip": "10.1.2.207",
"h2gcroc_port": 11000,
"pc_cmd_port": 11000,
"pc_data_port": 11001
}
}

## Slide 19: H2GCalib_3B – Socket Pool
python H2GCalib_3B/ 000_SocketPool.py
This program maintains multiple network connections, and other calibration programs use this as a relay, keep this program running in the background
[Pool] Control → 127.0.0.1:6002
[Pool] Data → 127.0.0.1:6001
...

## Slide 20: H2GCalib_3B – IO delay
python 007_IODelay.py
You can also use H2GDAQ’s IO Delay button to do this, but it don’t generate report

## Slide 21: H2GCalib_3B – ADC pedestal calibration
This program try to get the best pedestal value, it start from config “default_2024Aug_config.json”
python 005_PedestalCalib.py
Before calibration ADC pedestal
After calibration ADC pedestal
avg
target

## Slide 22: H2GCalib_3B – ToA Calibration
python 006_ToACalib.py -t 50 -i dump/005_PedestalCalib_data_20250625_135655/config_pede_a0_20250625_135817.json,dump/005_PedestalCalib_data_20250625_135655/config_pede_a1_20250625_135817.json
Before calibration ToA
After calibration ToA

## Slide 23: H2GCalib_3B – ToT Calibration
python 010_ToTCalib.py -t 300 -i dump/006_ToACalib_data_20250625_143708/config_toa_a0_20250625_151028.json,dump/006_ToACalib_data_20250625_143708/config_toa_a1_20250625_151028.json
Before calibration ToT
After calibration ToT

## Slide 24: H2GCalib_3B – inject 2V5 signal
modify file , line 282
python 009_IntInjection_2V5.py -i dump/010_ToTCalib_data_20250625_152005/config_tot_a0_20250625_155055.json,dump/010_ToTCalib_data_20250625_152005/config_tot_a1_20250625_155055.json
ASIC 0 (ch 1 ~ 9)
ASIC 1 is fine (ch 77 ~ 85)
phase_settings = [0, 1, 2, 3, 5, 8, 9, 10, 11, 12, 13, 14, 15]

## Slide 25: Operation Procedure - DAQ
Directory to store data
DAQ stop condition: By event count, by time or manual
Start/Stop DAQ button
high voltage relay setting
A0 = ASIC 0 only
A1 = ASIC1 only
F0 = ASIC0 and ASIC1 both
Send command to turn On/Off the relay base on the settings above. You also need to turn on your power supply
1
6
5
Output Run040.h2g

## Slide 26: Operation Procedure - Config
H2GConfig
  The GUI can read the configuration file and display it graphically. It allow you to modify register bit by bit
  It can upload configuration to the ASIC
  There is also a Ping FPGA button for you to test the ethernet connection
Test network connection
Load configuration
Upload to ASIC
3
2
4

## Slide 27: Operation Procedure - online_monitoring
https://github.com/tlprotzman/h2g_online_monitoring
  dependence: ROOT
  developed by Tristan Protzman from Lehigh University
$ git clone https://github.com/tlprotzman/h2g_online_monitoring.git
$ cd ~/workspace/h2g_online_monitoring
$ make
$ export DATA_DIRECTORY=~/workspace/H2GDAQ
$ export MONITORING_PORT=12345
$ root -q -b -x -l RunMonitoring.c\(RunNumber\)
Processing RunMonitoring.c(9)...
Parsing config file eeemcal_20sample.cfg
Key: MAX_SAMPLES, Value: 20
Key: NUM_FPGA, Value: 4
Key: DETECTOR_ID, Value: 2
Configuring for detector ID 2: EEEMCal
Configuration values:
NUM_LINES: 5
NUM_FPGA: 4
NUM_CHANNELS: 72
NUM_LINES: 5
MAX_SAMPLES: 20
MACHINE_GUN_MAX_TIME: 600
PACKET_SIZE: 1452
Configuring for detector ID 2: EEEMCal
Configuration values:
NUM_LINES: 5
NUM_FPGA: 4
NUM_CHANNELS: 72
NUM_LINES: 5
MAX_SAMPLES: 20
MACHINE_GUN_MAX_TIME: 600
PACKET_SIZE: 1452
huh??
Info in <THttpEngine::Create>: Starting HTTP server on port 12345
Attempting to open file /home/shyao/workspace/H2GDAQ/Run009.h2g
Starting at byte 828
Building events... done!
Updating canvases... done!
.....
$ cat eeemcal_20sample.cfg
MAX_SAMPLES=20
NUM_FPGA=4
DETECTOR_ID=2

## Slide 28: Operation Procedure - online_monitoring
Launch web browser and connect to localhost:12345
before calibration
After calibration

## Slide 29: Operation Procedure - online_monitoring
Inject light to crystal, each crystal use 4x4 SiPM. The light comes from an LED with a pulse width of 100ns
Light inject to this 16 channels
Light leaking to other crystal, must improve isolation
Crosstalk and Light leaking

## Slide 30: Operation Procedure - online_monitoring
Draw in crystal Array, it depend on your channel mapping
Light inject to this crystal
Light leaking to other crystal, must improve isolation
Other channel is not connected during test

## Slide 31: Operation Procedure - online_monitoring
Inject light to crystal, without leaking
Light inject to this 16 channels
Crosstalk (in H2GCROC)

## Slide 32: Operation Procedure - decode
https://github.com/tlprotzman/h2g_decode
  dependent ROOT , developped by Tristan Protzman from Lehigh University
  It provide a library h2g_decode and a executable h2g_run
$ cd ~/workspace
$ git clone https://github.com/tlprotzman/h2g_decode.git
$ cd ~/workspace/h2g_decode
$ mkdir data build; cd build; cmake ..; make
$ export DATA_DIRECTORY=~/workspace/H2GDAQ
$ export OUTPUT_DIRECTORY=~/workspace/h2g_decode
$ ./build/h2g_run
Usage: h2g_decode -r <run_number> [-d <detector_id>] [-n <num_kcu>] [-g] [-G LEVEL] [-T]
-r, --run Run number (required)
-d, --detector Detector ID (default: 0, LFHCAL: 1, EEEMCAL: 2)
-n, --num-kcu Number of KCUs (default: 4)
-g, --debug Enable debug output with INFO level
-G, --debug-level Set debug level explicitly:
0: OFF, 1: ERROR, 2: WARNING, 3: INFO, 4: DEBUG, 5: TRACE
-T, --truncate Enable ADC truncation
-h, --help Show this help message
** Data file is named with “Run%03d.h2g”

## Slide 33: Operation Procedure - decode
Convert rawdata to root file for future analysis.
$ ./build/h2g_run -r 14 -d 0 -n 1
$ root
root [1] TFile f(“Run014.root”)
root [2] events->Print()
******************************************************************************
*Tree :events : Events *
*Entries : 98258 : Total = 7360953537 bytes File Size = 492028194 *
* : : Tree compression factor = 14.96 *
******************************************************************************
*Br 0 :event_number : event_number/i *
*Entries : 98258 : Total Size= 395183 bytes File Size = 138994 *
*Baskets : 18 : Basket Size= 51200 bytes Compression= 2.84 *
*............................................................................*
*Br 1 :timestamps : timestamps[1]/i *
*Entries : 98258 : Total Size= 395145 bytes File Size = 114179 *
*Baskets : 18 : Basket Size= 51200 bytes Compression= 3.45 *
*............................................................................*
*Br 2 :num_samples : num_samples/i *
*Entries : 98258 : Total Size= 395161 bytes File Size = 4019 *
*Baskets : 18 : Basket Size= 51200 bytes Compression= 98.15 *
*............................................................................*
*Br 3 :adc : adc[144][32]/i *
*Entries : 98258 : Total Size= 1811620311 bytes File Size = 408563585 *
*Baskets : 5745 : Basket Size= 7088128 bytes Compression= 4.43 *
*............................................................................*
*Br 4 :toa : toa[144][32]/i *
*Entries : 98258 : Total Size= 1811620311 bytes File Size = 22647543 *
*Baskets : 5745 : Basket Size= 7088128 bytes Compression= 79.99 *
*............................................................................*
*Br 5 :tot : tot[144][32]/i *
*Entries : 98258 : Total Size= 1811620311 bytes File Size = 8514786 *
*Baskets : 5745 : Basket Size= 7088128 bytes Compression= 212.75 *
*............................................................................*
*Br 6 :hamming : hamming[144][32]/i *
*Entries : 98258 : Total Size= 1811643307 bytes File Size = 8529147 *
*Baskets : 5745 : Basket Size= 7089664 bytes Compression= 212.39 *
*............................................................................*
*Br 7 :hit_max : hit_max[144]/i *
*Entries : 98258 : Total Size= 56630819 bytes File Size = 21343750 *
*Baskets : 353 : Basket Size= 221184 bytes Compression= 2.65 *
*............................................................................*
*Br 8 :hit_pedestal : hit_pedestal[144]/i *
*Entries : 98258 : Total Size= 56632604 bytes File Size = 22016214 *
*Baskets : 353 : Basket Size= 221696 bytes Compression= 2.57 *
*............................................................................*
ADC waveform
x-axis: time
y-axis: amplitude

## Slide 34: Q/A
WWW

## Slide 35: Operation Procedure - online_monitoring
To modify channel mapping: online_monitor.cxx
// EEEMCal mapping - instead of "layers", we have a single plane, where each crystal is one connector
// FPGA IP | ID
// 208 | 0
// 209 | 1
// 210 | 2
// 211 | 3
int eeemcal_fpga_map[25] = {0, 0, 0, 0, 2,
0, 0, 0, 0, 2,
1, 1, 1, 1, 2,
1, 1, 1, 1, 2,
2, 2, 2, 2, 3};
// ASIC | ID
// 0 | 0
// 1 | 1
int eeemcal_asic_map[25] = { 0, 0, 0, 0, 1,
1, 1, 1, 1, 1,
0, 0, 0, 0, 1,
1, 1, 1, 1, 1,
0, 0, 0, 0, 0};
// Connector | ID
// A | 0
// B | 1
// C | 2
// D | 3
int eeemcal_connector_map[25] = { 0, 1, 2, 3, 0,
0, 1, 2, 3, 1,
0, 1, 2, 3, 2,
0, 1, 2, 3, 3,
0, 1, 2, 3, 0};
int eeemcal_16i_channel_a_map[16] = { 0, 1, 2, 3, 4, 5, 6, 7,
9, 10, 11, 12, 13, 14, 15, 16};
int eeemcal_16i_channel_b_map[16] = {19, 20, 21, 22, 23, 24, 25, 26,
27, 28, 29, 30, 31, 32, 33, 34};
int eeemcal_16i_channel_c_map[16] = {55, 56, 57, 58, 59, 60, 61, 62,
63, 64, 65, 66, 67, 68, 69, 70};
int eeemcal_16i_channel_d_map[16] = {36, 37, 38, 39, 40, 41, 42, 43,
45, 46, 47, 48, 49, 50, 51, 52};
/*
int eeemcal_16i_channel_a_map[16] = { 2, 6, 11, 15, 0, 4, 9, 13,
1, 5, 10, 14, 3, 7, 12, 16};
int eeemcal_16i_channel_b_map[16] = {20, 24, 29, 33, 18, 22, 27, 31,
19, 23, 28, 32, 21, 25, 30, 34};
int eeemcal_16i_channel_c_map[16] = {67, 63, 59, 55, 69, 65, 61, 57,
70, 66, 60, 56, 68, 64, 58, 54};
int eeemcal_16i_channel_d_map[16] = {50, 46, 40, 36, 52, 48, 42, 38,
51, 47, 43, 39, 49, 45, 41, 37};
*/

## Slide 36: Trigger path
ADC1 = ADC0 – ADC_pedestal
ADC2 = (ADC1 < ADC_TH) ? 0: ADC1
  ***if set ADC_pedestal = ADC_TH = 0
  ***ADC2 = ADC0
TOT1 = (TOT0 < TOT_THn) ? (TOT_THn – TOT_Pn) : (TOT0 – TOT_Pn)
  *** if set TOT_THn = TOT_Pn = 0
  *** TOT1 = TOT0
Charge = (TOT0 != 0) ? TOT1 * MultFactor : ADC2
12 bit
4 bit
7 bit
8 bit
8 bit
5 bit
10 bit

## Slide 37: Trigger path
ADC [ADC counts] is 10 bit
TOT [TOT counts] is 12 bit
linearization
Ref. TIPP21_HGCROC.pdf
MultFactor
TOT_
TOT_THn

## Slide 38: Local R&D
WWW

## Slide 39
Adapter board
Trigger board back
Trigger board front
Clock board back
Clock board front
Trigger board power
Clock board power

## Slide 40: Hardware @NCU
Device
  PbWO/LYSO Crystal 1/1
  SiPM board 1/1
  Adapter board 2/4
  protoboard 3/4
  KCU105 1/2
  Clock board 1/1
  Trigger board 1/1
Equip
  PC 1/1
  Power supply (for Clock distribution board and Trigger board) 1/1
  Signal generator 1/1
  Network Switch 1GbE 1/1
Cable
  SMA male to male cable 3/8
  SMA to BNC connector 2/6
  SMA to <unknown> cable for clock board connect to <unknown input> 0/1
  SMA to <unknown> cable for trigger board connect to signal generator 0/1
  Power cable for Clock distribution board 1/1
  Power cable for Trigger board 1/1
  SiPM board to Adapter board ?/? not sure
  Network Cable 1GbE 1/2
Assume 2 KCU105 and 4 protoboard setup

## Slide 41: Current environment
Waveform generator
Power supply
Adapter board
SiPM and Crystal
Clock
Clock
Oscilloscope
Splitter cable

## Slide 42: About 12V Power
GND | NC | 12V
GND | NC | 12V
Shared FPGA evaluation board power supply via splitter
GND | NC | 12V
GND | NC | 12V
GND | GND | NC
NC | 12V | 12V
NC | NC | NC
GND | NC | 12V
KCU105 Power supply
KCU 105
Trigger board
Clock board
12V
GND
12V
GND
12V
GND
12V
GND
 Connect all them together, your cable will be
GND | GND | 12V
GND | 12V | 12V

## Slide 43: Clock board
.
7 N
Jumper select:
Int: On board 40MHz 0.5ppm clock
Ext: External input
40MHz
7 P
6 N
6 P
0 N
1 P
0 P
1 N
2 P
2 N
5 N
5 P
4 P
4 N
3 N
3 P
The output of 7P and 6P

## Slide 44: external trigger signal
1.8V output, 10Hz, Pulse width 25ns
Send to Trigger input
Input
Output
If the input signal voltage decreases, the output width will decrease. Smaller than 0.8V output disappears
1MΩ

## Slide 45: UDP packet format (FPGA -> PC)
a UDP packet
1452 or 8972 Bytes
2B | Byte Counter | (0 ~ 36) *40
2B | Packet Counter | 0 ~ 65535
8B | String (“0”) | 
Normal frame N = 0 ~ 36
JUMBO frame N = 0 ~ 224
1B | Header | b0-3: Asic address. b4-7: 0xA0
1B | Fpga board address | 0x00 ~ 0x07
1B | packet type | TR0: 0x20, TR1: 0x21, TR2: 0x22, TR3: 0x23, DQ0: 0x24, DQ1: 0x25
1B | packet id | 0 ~ 4
Header | Timestamp | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B)
Ethernet, IP, UDP Header | UDP Header (12B) | Data (40B) | Data (40B) | Data (40B) | ... | ... | ... |  | Data (N * 40B)
Data format header:
UDP header:
Data (40B):
32B | Timestamp | Guess: Unit 25ns, for time sync between FPGAs. wrapping at 0x40000000, 26.8435456 s
Timestamp:
V4.11

## Slide 46: UDP packet format (FPGA -> PC)
One read command generate ten of 40B packet (400Bytes), Two ASIC is 800 Bytes
Timestamp field is duplicated
0, 0, 0x24, 0 | Timestamp | DaqH | CM | CH0 | CH1 | CH2 | CH3 | CH4 | CH5
0, 0, 0x24, 1 | Timestamp | CH6 | CH7 | CH8 | CH9 | CH10 | CH11 | CH12 | CH13
0, 0, 0x24, 2 | Timestamp | CH14 | CH15 | CH16 | CH17 | Calib | CH18 | CH19 | CH20
0, 0, 0x24, 3 | Timestamp | CH21 | CH22 | CH23 | CH24 | CH25 | CH26 | CH27 | CH28
0, 0, 0x24, 4 | Timestamp | CH29 | CH30 | CH31 | CH32 | CH33 | CH34 | CH35 | CRC-32
0, 0, 0x25, 0 | Timestamp | DaqH | CM | CH0 | CH1 | CH2 | CH3 | CH4 | CH5
0, 0, 0x25, 1 | Timestamp | CH6 | CH7 | CH8 | CH9 | CH10 | CH11 | CH12 | CH13
0, 0, 0x25, 2 | Timestamp | CH14 | CH15 | CH16 | CH17 | Calib | CH18 | CH19 | CH20
0, 0, 0x25, 3 | Timestamp | CH21 | CH22 | CH23 | CH24 | CH25 | CH26 | CH27 | CH28
0, 0, 0x25, 4 | Timestamp | CH29 | CH30 | CH31 | CH32 | CH33 | CH34 | CH35 | CRC-32
Possibly crosses UDP packet boundaries
DQ0, Half H2GCROC
DQ1, Half H2GCROC
A DAQ frame (Next page)
(fpga id, asic id, packet type, packet id)
Header | Timestamp | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B) | Data (4B)
V4.11

## Slide 47: DAQ frame (H2GCROC->FPGA)
A read command, send a DAQ frame (From datasheet)
40 * 32bit = 1280 bits
DaqH | CM | CH0 | CH1 | ... | CH17 | Calib | CH18 | ... | CH35 | CRC-32 | IDLE
4b | hd | 0b1111
12b | BxCounter | Value of the Bunch Crossing Counter (BCD) on 12 bits. (40MHz). Range: 1 - 3564. Reset value: 1
6b | Event | Value of the Event Counter (EC) , to detect if a L1 trigger has been sent but related data wasn't saved (RAM2 full). Range: 0 - 63. Reset value: 1
3b | Orbit | Value of the Orbit Counter (OC). Range: 0 - 7. Reset value: 0. Increase 1 when BCD wrapping
1b | H1 | Error during hamming decoding in Header
1b | H2 | Error during hamming decoding in channels from CM to CH17 (1st Quarter)
1b | H3 | Error during hamming decoding in channels from Calib to CH35 (2nd Quarter)
4b | tr | 0b0101: no first event flag 0b0010: first event flag
DaqH:
10b | CM1 ADC
10b | CM0 ADC
12b | 0b000
CM:
 | CASE1, CASE2 | CASE3 | CASE4
10b | TOA | TOA | TOA
10b | ADC | TOT | TOT
10b | ADC-1 | ADC-1 | ADC
1b | Tp | Tp | Tp
1b | 0 | 1 | Tc
CH#, Calib:
~ 89us BxCounter wrapping
~ 712us OC wrapping
41 x 32bits
V4.11
CASE1: Charge Q < TOA Threshold, TOA = 0
CASE2: Charge Q < TOT Threshold, TOA > 0
CASE3: Charge Q > TOT Threshold
CASE4: Debug mode
Tp: TOT-In-Progress
Tc: TOT Complete
-1: Previous BX ?

## Slide 48: Decode
UDP_TX_Packet_Counter: 57 ByteCounter: 1440
hd: 15 H1: 0 H2: 0 H3: 0 tr: 5 BxCounter: 2827 Event: 13 Orbit: 0
FPGA: 0 ASIC: 0 packet_type: 25 packet_id: 00 timestamp: 0ad3daca Data: fb0b3405 00017000 06800000 09600000 08b00000 06600000 07800000 07800000
FPGA: 0 ASIC: 0 packet_type: 25 packet_id: 01 timestamp: 0ad3daca Data: 06500000 07100000 06200000 06800000 07900000 05f00000 06200000 07900000
FPGA: 0 ASIC: 0 packet_type: 25 packet_id: 02 timestamp: 0ad3daca Data: 08000000 05a00000 04600000 07a00000 05800000 06f00000 07a00000 07b00000
FPGA: 0 ASIC: 0 packet_type: 25 packet_id: 03 timestamp: 0ad3daca Data: 09e00000 06e00000 06700000 09200000 06f00000 05500000 09300000 07200000
FPGA: 0 ASIC: 0 packet_type: 25 packet_id: 04 timestamp: 0ad3daca Data: 07000000 08500000 05b00000 07d00000 06e00000 08b00000 05100000 ec0d330c
timestamp differ 181656266
hd: 15 H1: 0 H2: 0 H3: 0 tr: 5 BxCounter: 3130 Event: 13 Orbit: 7
FPGA: 0 ASIC: 1 packet_type: 24 packet_id: 00 timestamp: 0ad3daca Data: fc3a3785 0001c400 09200000 03b00000 07200000 05b00000 08100000 05c00000
FPGA: 0 ASIC: 1 packet_type: 24 packet_id: 01 timestamp: 0ad3daca Data: 06900000 08500000 08e00000 08400000 08200000 08600000 08000000 0830019b
FPGA: 0 ASIC: 1 packet_type: 24 packet_id: 02 timestamp: 0ad3daca Data: 09800000 08100000 09600000 08b00000 05900000 09000000 07800000 07d00000
FPGA: 0 ASIC: 1 packet_type: 24 packet_id: 03 timestamp: 0ad3daca Data: 05f00000 08500000 08600000 08c00000 06b00000 07800000 07400000 07900000
FPGA: 0 ASIC: 1 packet_type: 24 packet_id: 04 timestamp: 0ad3daca Data: 08900000 07200000 07b00000 08100000 07100000 07100000 06200000 5133d9f4
hd: 15 H1: 0 H2: 0 H3: 0 tr: 5 BxCounter: 2827 Event: 13 Orbit: 0
FPGA: 0 ASIC: 0 packet_type: 24 packet_id: 00 timestamp: 0ad3daca Data: fb0b3405 00013000 06100000 06400000 08d00000 07d00000 0a500000 06000000
FPGA: 0 ASIC: 0 packet_type: 24 packet_id: 01 timestamp: 0ad3daca Data: 06400000 07900000 05900000 07800000 03f00000 07c00000 04900000 07f00000
FPGA: 0 ASIC: 0 packet_type: 24 packet_id: 02 timestamp: 0ad3daca Data: 04e00000 06600000 05700000 05700000 05600000 03100000 04a00000 03200000
FPGA: 0 ASIC: 0 packet_type: 24 packet_id: 03 timestamp: 0ad3daca Data: 03100000 04700000 05200000 06300000 05800000 03e00000 04e00000 05a00000
FPGA: 0 ASIC: 0 packet_type: 24 packet_id: 04 timestamp: 0ad3daca Data: 06000000 07600000 3ff00000 06b00000 05e00000 05800000 04400000 ec98fb1f
hd: 15 H1: 0 H2: 0 H3: 0 tr: 5 BxCounter: 3130 Event: 13 Orbit: 7
FPGA: 0 ASIC: 1 packet_type: 25 packet_id: 00 timestamp: 0ad3daca Data: fc3a3785 00017800 06a00000 05500000 03600000 05f00000 06100000 07000000
FPGA: 0 ASIC: 1 packet_type: 25 packet_id: 01 timestamp: 0ad3daca Data: 05d00000 05a00000 06d00000 05300000 06d00000 06000000 05600000 04900000
FPGA: 0 ASIC: 1 packet_type: 25 packet_id: 02 timestamp: 0ad3daca Data: 04f00000 08000000 06900000 06400000 07000000 04600000 07500000 05600000
FPGA: 0 ASIC: 1 packet_type: 25 packet_id: 03 timestamp: 0ad3daca Data: 07600000 06300000 03f00000 05600000 05c00000 05c00000 04200000 03a00000
FPGA: 0 ASIC: 1 packet_type: 25 packet_id: 04 timestamp: 0ad3daca Data: 05d00000 08e00000 07a00000 01900000 06a00000 02b00000 05a00000 dedc3d7c
...
V4.11

## Slide 49: Raw Data
A custom program can send start_daq and store UDP packet
red boxed fields require byte swapping, i.e. Big-endian
[shyao@cats myprj]$ od -Ax -t x4 -N 12 data/rawdata.10.bin
000000 5300a005 00000000 00000000
00000c
[shyao@cats myprj]$ od -Ax -t x4 -j 12 -w40 -N 1452 data/rawdata.10.bin
00000c 002400a0 9c626808 85ad3cf5 00280100 00009005 0000d005 00006008 00001001 0000d009 0000b005
000034 012400a0 9c626808 0000a005 0000a00f 0000b005 00004007 0000b003 00005007 0000f003 00008007
00005c 022400a0 9c626808 00006004 00002006 0000f03f 0000f03f 00004005 0000f03f 00002004 0000e002
000084 032400a0 9c626808 00008002 00005004 0000a004 0000f005 00002005 00009003 0000c004 00005005
0000ac 042400a0 9c626808 0000b005 0000f03f 0000f03f 00005006 00006005 0000f03f 0000d003 cb3a5347
0000d4 002500a1 9c626808 05ad6bf6 00780100 0000c006 00005005 00008003 0000e005 0000c005 00002007
0000fc 012500a1 9c626808 0000f005 0000a005 0000d006 00004005 0000e006 00001006 00009005 00009004
000124 022500a1 9c626808 0000f004 00004008 0000a006 00002006 00003007 00005004 00002007 00006005
00014c 032500a1 9c626808 00008007 00006006 00002004 0000a005 0000d005 0000d005 00001004 0000d003
000174 042500a1 9c626808 0000d005 00002009 0000a007 00009001 0000a006 0000d002 0000e005 da35dee7
00019c 002500a0 9c626808 85ad3cf5 41010400 00001010 00000000 00009000 00001000 00000010 00001000
0001c4 012500a0 9c626808 00001000 00000000 00001000 00000000 00001000 00009017 00001000 00006015
0001ec 022500a0 9c626808 00009018 00001015 00000000 00001000 00001010 00001000 00001010 0000f017
000214 032500a0 9c626808 00000000 00003018 00001000 00001001 00000000 00000000 00000019 00001010
00023c 042500a0 9c626808 0000d016 00001010 00000010 00008019 fc010000 00001019 00002010 625b12cb
000264 002400a1 9c626808 05ad6bf6 00c80100 00000009 0000d003 00002007 0000a005 00002008 00009005
00028c 012400a1 9c626808 0000a006 00004008 0000f008 00006008 00001008 00008008 0000f007 00004008
0002b4 022400a1 9c626808 00006009 00000008 00003009 0000a008 00007005 0000f008 00008007 0000b007
0002dc 032400a1 9c626808 00000006 00003008 00006008 00008008 0000a006 00007007 00005007 00009007
000304 042400a1 9c626808 0000a008 00001007 0000d007 0000f007 0000f006 0000f006 0000f005 ce6c3496
00032c 002400a0 c5626808 85b13df5 00300100 0000a005 0000c005 00008008 00001001 0000c009 0000b005
000354 012400a0 c5626808 0000c005 0000600f 0000a005 00004007 0000b003 00005007 00000004 00006007
00037c 022400a0 c5626808 00008004 00002006 0000f03f 0000f03f 00002005 0000f03f 00006004 0000f002
0003a4 032400a0 c5626808 0000b002 00004004 0000c004 0000e005 00003005 0000a003 0000f004 00008005
0003cc 042400a0 c5626808 0000e005 0000f03f 0000f03f 00005006 00006005 0000f03f 0000f003 808acca5
0003f4 002500a1 c5626808 05b16cf6 00780100 0000d006 00004005 00009003 0000e005 0000e005 00001007
00041c 012500a1 c5626808 00000006 00009005 0000d006 00004005 0000f006 00001006 0000a005 0000b004
000444 022500a1 c5626808 00000005 00004008 0000a006 00005006 00002007 00004004 00005007 00006005
00046c 032500a1 c5626808 00008007 00008006 0000f003 0000a005 0000f005 0000e005 0000f003 0000d003
000494 042500a1 c5626808 0000d005 00000009 00009007 00008001 0000c006 0000d002 0000d005 f1be7c7b
0004bc 002500a0 c5626808 85b13df5 41010400 00001010 00000000 00009000 00001000 00000010 00001000
0004e4 012500a0 c5626808 00001000 00000000 00001000 00000000 00001000 00009017 00001000 00006015
00050c 022500a0 c5626808 00009018 00001015 00000000 00001000 00001010 00001000 00001010 0000f017
000534 032500a0 c5626808 00000000 00003018 00001000 00001001 00000000 00000000 00000019 00001010
00055c 042500a0 c5626808 0000d016 00001010 00000010 00008019 00000000 00001019 00002010 70dc1c4d
000584 002400a1 c5626808 05b16cf6 00c40100 00001009 0000d003 00001007 0000b005 00002008 0000c005
0005ac 5400a005 00000000 00000000
0005b8
------ 00000000 00000000 00000000 00000000 00000000 00000000 00000000 00000000 00000000 00000000
UDP Header
Next UDP Header
V4.11
For firmware V4.11, the last UDP packet is truncated. And it also send a full zero packet when stopping DAQ.

## Slide 50: DAQ mode
In practice, the external trigger signal is sent to the FPGA.
Four possible mode
  DAQ Push
  Generator
  Internal trigger
  External trigger
Common parameter - Machine Gun
  The FPGA firmware generates a series of read commands and sends them to H2GCROC.
  The real number of fast read commands sent will be the setting value plus 1.
  This field has 8 digits, but the reasonable value range is 1 to 31.
V4.11

## Slide 51: DAQ Push
Software trigger
Send this trigger by user interface
V4.11

## Slide 52: External trigger
test signal : 1MHz, 10 pulse
V4.11
test signal : 32KHz, 1000 pulse

## Slide 53: Generator
Generator Interval + 1
The generator function is enabled by user commands and is executed once after being enabled.
Three parameters
  Pre Interval – delay before the first fire, 16 bit, Unit 25ns
  Generator Cycle - fire how many time, 32 bit
  Generator Interval – Time between two fires, 32 bit, Unit 25ns
Number of fast read command send = (Generator Cycle) * (Machine Gun + 1)
Generator Cycle
Machine Gun + 1
Pre Interval
t
Notes: Simply say, let the FPGA generates periodic read signals for testing.

## Slide 54: Internal trigger
No way was found to enable this feature from the user interface. Try using custom code.
FPGA firmware manual say it only support ASIC 0

## Slide 55: Data rate / trigger rate
Each channel generates 4 bytes of data (include TOA, TOT, ADC) every 25 nanoseconds.
A DAQ Frame is 1280+32 bits, 38 channels, it takes 1.025 microseconds to transmit data to FPGA.
If maximum consecutively read (Machine Gun = 31, 800ns window)
  H2GCROC takes 32.8 microseconds to transmit data to FPGA.
  FPGA takes 163.8 204.8 microseconds to transmit data to PC.
  Maximum average trigger rate is about 6 5kHz
FPGA 0
Network Switch 0
PC
Data transfer in parallel, Each lane handle half channels. ~5 Gbps
ASIC 0
1.28Gbps
ASIC 1
1 Gbps
1 Gbps
1.28Gbps
1.28Gbps
1.28Gbps
Buffer
Buffer
Buffer
Buffer
Buffer
Buffer
Buffer
Other FPGAs
Other PCs
1312 bits = header + half channel + crc + idle
Trigger
Any bottleneck will lead to longer dead time or potential data loss.

## Slide 56: DAQ test
Set generator parameter to 1KHz rate, 100k cycles. Set machine gun to 32 read
Set High voltage to 28V.
All other connected 63 channels is similar.
ADC waveform
x-axis: time
y-axis: amplitude

## Slide 57: TOT
TOT is 12 bit in ASIC, but send 10 bit only
0 | b8 | b7 | b6 | b5 | b4 | b3 | b2 | b1 | b0
b11 | b10 | b9 | b8 | b7 | b6 | b5 | b4 | b3 | b2 | b1 | b0
1 | b11 | b10 | b9 | b8 | b7 | b6 | b5 | b4 | b3
if all are zero
Otherwise send:
In ASIC:
decode
Send:

## Slide 58: TOA
TOA output is not zero when signal over the threshold
  I use it as the filter
28V, Efficiency: 4736/32/98258 = ~ 0.1% ~ 0.3%

## Slide 59: Noise
28V. Event Filter: (TOA == 0)
pedestal calibration is wrong
28V Noise

## Slide 60: Noise
Event Filter: (TOA == 0)
float
0V
28V
10V

## Slide 61: Count
28V

## Slide 62: ProjectionY
28V. Event Filter: (TOA > 0)
All channel there is a peak.

## Slide 63: LDE TEST layout
-
+
Switch
PC
1GbE Ethernet
1GbE Ethernet
Trigger board
Clock board
Power cable splitter
LED signal
4P PCB3.81
端子台
KCU105 Power adapter 12V
Adapter board (frontside)
Adapter board (backside)
KCU105 10.1.2.208
Protoboard
SMA to SMA
SMA to SMA
SMA to SMA
SiPM Array & Led in dark box
SiPM Bias voltage 25 ~ 32V
ASIC0
ASIC1
waveform generator
KEYSIGHT 33520B
10.1.2.2
Power supply
GWINSTEK GPP-3060
10.1.2.30
Ch 2
Ch 1
(BNC to SMA)
to scope
SCPI programming
Ch 1
Almalinux 9.5
10.1.2.207
63
0
0
0
70
70

## Slide 64: LED test signal
Unknown delay between Led light and Trigger
Tune signal delay or firmware parameter
Ch2: To Led
Ch1: Trigger
Signal delay (Tunable)
Amp (Tunable)

## Slide 65: LED test result
ADC waveform. Use online_monitor
run27
signal delay 0
Bias 29V
LED 1.3V, 32ns
Ext trig. delay 10
run28
signal delay 0
Bias 29V
LED 1.3V, 32ns
Ext trig. delay 0

## Slide 66: LED test result
run29
signal delay 0
Bias 29V
LED 1.5V, 32ns
Ext trig. delay 10
run30
signal delay 0
Bias 29V
LED 1.4V, 32ns
Ext trig. delay 0

## Slide 67: LED test result
run31
signal delay 0
Bias 29V
LED 1.4V, 25ns
Ext trig. delay 10
run32
signal delay 0
Bias 29V
LED 1.45V, 25ns
Ext trig. delay 10

## Slide 68: LED test result
run33
signal delay 0
Bias 28V
LED 1.4V, 25ns
Ext trig. delay 10
run34
signal delay 0
Bias 30V
LED 1.45V, 25ns
Ext trig. delay 10

## Slide 69: Channel mapping
Multiple stage....
  SiPM board layout
  Cable reverse order
  Protoboard layout
  ASIC 0/1?
  FPGA 0
Lookup table
  https://docs.google.com/spreadsheets/d/1VDs9stUrkrNTq3McJMH1bralzvcX4AuQ8HcpJwmnNY8/edit?usp=drive_link

## Slide 70: More
Run 32, ASIC1, CH10 ADC waveform
Is this jitter or photon count?
0
histogram
300
adc waveform
0
20

## Slide 71: decoder broken
ASIC 0 disable?
$ ./build/h2g_run -r 28 -d 0 -n 1
$ root
root [1] TFile f(“Run28.root”)
root [2] events->Print()
******************************************************************************
*Tree :events : Events *
*Entries : 0 : Total = 8669 bytes File Size = 1057 *
* : : Tree compression factor = 1.00 *
******************************************************************************
*Br 0 :event_number : event_number/i *
*Entries : 0 : Total Size= 513 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 1 :timestamps : timestamps[1]/i *
*Entries : 0 : Total Size= 511 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 2 :num_samples : num_samples/i *
*Entries : 0 : Total Size= 509 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 3 :adc : adc[144][32]/i *
*Entries : 0 : Total Size= 495 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 4 :toa : toa[144][32]/i *
*Entries : 0 : Total Size= 495 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 5 :tot : tot[144][32]/i *
*Entries : 0 : Total Size= 495 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
*Br 6 :hamming : hamming[144][32]/i *
*Entries : 0 : Total Size= 511 bytes One basket in memory *
*Baskets : 0 : Basket Size= 32000 bytes Compression= 1.00 *
*............................................................................*
...
ADC waveform
x-axis: time
y-axis: amplitude
?

## Slide 72: SCPI programming
pip install pyvisa pyvisa-py
import pyvisa
rm = pyvisa.ResourceManager()
#KEYSIGHT 33520B waveform generator
wavegen = rm.open_resource('TCPIP0::10.1.2.2::inst0::INSTR')
wavegen.read_termination = '\n'
wavegen.write_termination = '\n'
wavegen.write('*RST')
wavegen.write('SOURCE1:FUNCTION PULSE')
wavegen.write('SOURCE1:FUNCTION:PULSE:TRANSITION:LEADING 8.4 ns') # instrument minimum is 8.4 ns
wavegen.write('SOURCE1:FUNCTION:PULSE:TRANSITION:TRAILING 8.4 ns')
wavegen.write('SOURCE1:FUNCTION:PULSE:WIDTH 25 ns') # instrument minimum is 16 ns
wavegen.write('SOURCE1:FREQUENCY 1000')
wavegen.write('SOURCE1:VOLTAGE:HIGH 1.3')
wavegen.write('SOURCE1:VOLTAGE:LOW 0')
wavegen.write('OUTPUT1 ON')
wavegen.write('SOURCE2:FUNCTION PULSE')
wavegen.write('SOURCE2:FUNCTION:PULSE:WIDTH 25 ns')
wavegen.write('SOURCE2:FUNCTION:PULSE:TRANSITION:LEADING 8.4 ns')
wavegen.write('SOURCE2:FUNCTION:PULSE:TRANSITION:TRAILING 8.4 ns')
wavegen.write('SOURCE2:FREQUENCY 1000')
wavegen.write('SOURCE2:VOLTAGE:HIGH 1.45') # voltage
wavegen.write('SOURCE2:VOLTAGE:LOW 0')
wavegen.write('OUTPUT2 ON')
# let both channel's pulse synchronized, this channel phase is independent to burst mode phase
wavegen.write('SOURCE2:PHASE:SYNCHRONIZE')
wavegen.write('SOURCE2:PHASE:REFERENCE')
wavegen.write('UNIT:ANGLE SECOND')
wavegen.write('SOURCE1:PHASE 0.0000003') # 300ns
print("ERROR: ", wavegen.query('SYSTEM:ERROR?'))
wavegen.close()
import pyvisa
rm = pyvisa.ResourceManager()
#GWINSTEK GPP-3060 Programmable DC Power Supply
ps = rm.open_resource('TCPIP0::10.1.2.30::1026::SOCKET')
ps.read_termination = '\n'
ps.write_termination = '\n'
ps.write('*RST')
ps.write(':VSET1:28') # CH1 voltage range = 0 ~ 32V
ps.write(':ISET1:1') # CH1 current range = 0 ~ 6.2A
ps.write(':OUTPUT1:STATE ON')
#ps.write(':ALLOUTOFF')
ps.close()

## Slide 73: Two FPGA
mechanism needed
  Is there enough space on the trigger board and clock board?
  Is the SMA to SMA cable long enough?
  space for 10G switch and network cable length?
Power cable for SiPM?
Prepare PC for beam test

## Slide 74: decoder
ASIC 0 must work
1KHz, 10000 trigger => 9821 event
2 FPGA * 2 ASIC = 288 channel
$ ./build/h2g_run -r 40 -d 0 -n 2
$ root
root [1] TFile f(“Run040.root”)
root [2] events->Print()
******************************************************************************
*Tree :events : Events *
*Entries : 9824 : Total = 1472807942 bytes File Size = 77986302 *
* : : Tree compression factor = 18.91 *
******************************************************************************
*Br 0 :event_number : event_number/i *
*Entries : 9824 : Total Size= 40052 bytes File Size = 14069 *
*Baskets : 3 : Basket Size= 51200 bytes Compression= 2.81 *
*............................................................................*
*Br 1 :timestamps : timestamps[2]/i *
*Entries : 9824 : Total Size= 79340 bytes File Size = 26918 *
*Baskets : 3 : Basket Size= 51200 bytes Compression= 2.93 *
*............................................................................*
*Br 2 :num_samples : num_samples/i *
*Entries : 9824 : Total Size= 40045 bytes File Size = 543 *
*Baskets : 3 : Basket Size= 51200 bytes Compression= 72.81 *
*............................................................................*
*Br 3 :adc : adc[288][32]/i *
*Entries : 9824 : Total Size= 362491179 bytes File Size = 65978440 *
*Baskets : 3684 : Basket Size= 7088640 bytes Compression= 5.49 *
*............................................................................*
*Br 4 :toa : toa[288][32]/i *
*Entries : 9824 : Total Size= 362491179 bytes File Size = 2064431 *
*Baskets : 3684 : Basket Size= 7088640 bytes Compression= 175.55 *
*............................................................................*
*Br 5 :tot : tot[288][32]/i *
*Entries : 9824 : Total Size= 362491179 bytes File Size = 2036343 *
*Baskets : 3684 : Basket Size= 7088640 bytes Compression= 177.97 *
*............................................................................*
*Br 6 :hamming : hamming[288][32]/i *
*Entries : 9824 : Total Size= 362505931 bytes File Size = 2466907 *
*Baskets : 3684 : Basket Size= 7089152 bytes Compression= 146.92 *
*............................................................................*
*Br 7 :hit_max : hit_max[288]/i *
*Entries : 9824 : Total Size= 11333891 bytes File Size = 2611281 *
*Baskets : 170 : Basket Size= 221696 bytes Compression= 4.34 *
*............................................................................*
*Br 8 :hit_pedestal : hit_pedestal[288]/i *
*Entries : 9824 : Total Size= 11334761 bytes File Size = 2682917 *
*Baskets : 170 : Basket Size= 221696 bytes Compression= 4.22 *
*............................................................................*
root [1] .q

## Slide 75: Discuss
Practical limitations
  LEDs are not ideal light sources; as point sources, their brightness varies with distance.
  Is it possible to use an integrating sphere, or is there a better method?
  The waveform generator's rising and falling edges are too slow.
  SIPM is affected by temperature and currently does not have any temperature control in laboratory.

## Slide 76: # Events vs Trigger frequency
 | 1KHZ | 2KHZ | 3KHZ | 4KHZ | 5KHZ | 6KHZ | 7KHZ | 8KHZ | 9KHZ | 10KHZ
4 ASIC Gun=31 | 9818 | 9494 | 6078 | 4471 | 87 crash | 61 crash | 44 crash |  |  | 
4 ASIC Gun=15 | 9915 | 9911 | 9911 | 9905 | 7624 | 6365 | 5483 | 4765 | 4311 | 
2 ASIC Gun=31 | 9822 | 9821 | 9823 | 9826 | 129 crash | 64 crash | 42 crash | 30 crash | 25 crash | 19 crash
2 ASIC Gun=15 | 9910 | 9917 | 9910 | 9902 | 9927 | 9919 | 9895 | 9930 | 8488 | 264 crash
10000 Trigger (from waveform generator)
4 ASIC = 2 KCU x 2 ASIC
2 ASIC = 1 KCU x 2 ASIC
Crash = decoder error
F0
Network Switch 0
PC
A0
1.28 Gbps
A1
1 Gbps
2.5 Gbps
1.28 Gbps
1.28 Gbps
1.28 Gbps
Trigger
F1
A0
1.28 Gbps
A1
1.28 Gbps
1.28 Gbps
1.28 Gbps
1 Gbps

## Slide 77: Trigger delay
ext_trigger_delay = 0
We use CITIROC as trigger source
  LED light -> SiPM -> 32 Ch to CITIROC -> trigger out
  -> 32 Ch to H2GCROC -> trigger in -> output sample data
No signal peak found
ext_trigger_delay = 32
ext_trigger_delay = 64

## Slide 78: Trigger delay
By tuning H2GCROC register “L1Offset” from 8 to 33, keep ext_trigger_delay = 0
L1offset = 8
L1offset = 0
L1offset = 32
L1offset = 33
L1offset = 48

## Slide 79: Trigger delay
H2GConfig

## Slide 80: event matching
The timestamp values ​​in the original root file have been rewrite
Add a new column “timestamp_r” and its value is the same as that in rawdata.
Work in progress.
******************************************************************************
*Tree :events : Events *
*Entries : 3353979 : Total = 158473711591 bytes File Size = 9461118948 *
* : : Tree compression factor = 16.75 *
******************************************************************************
*Br 0 :event_number : event_number/i *
*Entries : 3353979 : Total Size= 13450993 bytes File Size = 4711578 *
*Baskets : 344 : Basket Size= 51200 bytes Compression= 2.85 *
*............................................................................*
*Br 1 :timestamps : timestamps[1]/i *
*Entries : 3353979 : Total Size= 13450303 bytes File Size = 4665670 *
*Baskets : 344 : Basket Size= 51200 bytes Compression= 2.88 *
*............................................................................*
*Br 2 :timestamps_r : timestamps_r[1]/i *
*Entries : 3353979 : Total Size= 13450999 bytes File Size = 12137277 *
*Baskets : 344 : Basket Size= 51200 bytes Compression= 1.11 *
*............................................................................*
*Br 3 :num_samples : num_samples/i *
*Entries : 3353979 : Total Size= 13450645 bytes File Size = 105228 *
*Baskets : 344 : Basket Size= 51200 bytes Compression= 127.76 *
*............................................................................*
*Br 4 :adc : adc[144][20]/i *
*Entries :3353979 : Total Size=38638760143 bytes File Size = 7979992633 *
*Baskets : 10019 : Basket Size= 7802368 bytes Compression= 4.84 *
*............................................................................*
*Br 5 :toa : toa[144][20]/i *
*Entries :3353979 : Total Size=38638760143 bytes File Size = 174550604 *
*Baskets : 10019 : Basket Size= 7802368 bytes Compression= 221.36 *
*............................................................................*
*Br 6 :tot : tot[144][20]/i *
*Entries :3353979 : Total Size=38638760143 bytes File Size = 170801890 *
*Baskets : 10019 : Basket Size= 7802368 bytes Compression= 226.22 *
*............................................................................*
*Br 7 :hamming : hamming[144][20]/i *
*Entries :3353979 : Total Size=38638800235 bytes File Size = 177674724 *
*Baskets : 10019 : Basket Size= 7803392 bytes Compression= 217.47 *
*............................................................................*
*Br 8 :hit_max : hit_max[144]/i *
*Entries : 3353979 : Total Size= 1932400547 bytes File Size = 450786221 *
*Baskets : 5295 : Basket Size= 390144 bytes Compression= 4.29 *
*............................................................................*
*Br 9 :hit_pedestal : hit_pedestal[144]/i *
*Entries : 3353979 : Total Size= 1932427042 bytes File Size = 485242055 *
*Baskets : 5295 : Basket Size= 390144 bytes Compression= 3.98 *
*............................................................................*

## Slide 81: NA22 Test
Run 12, 10971 event, machine_gun = 31
  TOA > 0 cut, 5608 event left
ADC spectrum
Peak or Integrate (= AVERAGE / SAMPLEs)
Need more event

## Slide 82: NA22 Test
Run 14, 3353979 event, machine_gun = 19
  TOA > 0 cut, 593615 event left
ADC spectrum
Peak or Integrate (= AVERAGE / SAMPLEs)
This analysis result still needs further check.

## Slide 83: Na22 Test: ADC fitting by channel
Fitting function "gaus(0)+expo(3)" = [0]*exp(-0.5*((x-[1])/[2])**2) + exp([3]+[4]*x)
Run14, 1 minutes, 220000 events, 3.5K per channel
Gun = 20
Run14, 3353976 events, ~15.4 minutes
Average trigger rate: 3.6 KHz
Gun = 20

## Slide 84: Na22 Test: ADC vs Energy
Run30, 323337 evts, 68.5s, ATR: 4.7 KHz
SiPM 28V
Gun = 20
Run31, 318145 evts, 66.5s, ATR: 4.7 KHz
SiPM 28.5V
Gun = 20
Run33, 348422 evts, 74.8s, ATR: 4.6 KHz
SiPM 29V
Gun = 20

## Slide 85: Na22 Test: ADC vs Energy
Run34, 295130 evts, 64.1s, ATR: 4.6 KHz
SiPM 29.5V
Gun = 20
Run35, 304667 evts, 66.2s, ATR: 4.6 KHz
SiPM 30V
Gun = 20

## Slide 87: Voltage tuning
Reg | Voltage
0x00 | 2.3789
0x08 | 2.2996
0x10 | 2.0128
0x18 | 1.7966
0x20 | 1.4773
0x28 | 1.2021
0x30 | 0.9086
0x38 | 0.6403
0x3F | 0.5125
Chart: 
| 0 | 8 | 16 | 24 | 32 | 40 | 48 | 56 | 63
 | 2.3789 | 2.2996 | 2.0128 | 1.7966 | 1.4773 | 1.2021 | 0.9086 | 0.6403 | 0.5125

## Slide 89: Unknown problem
Both protoboard ASIC0, channel 36-71

## Slide 90: To be continue...