## Slide 1: ECal 3rd Prototype Beam Test Analysis
Yen-Chen Wu
Dept. Physics, National Central University
2026/09/09

## Slide 2: Energy Resolution Analysis Approach
ADC raw max
ADC max after fitting
ADC waveform integration
TOT
TOA
check resolution of E1x1, E3x3, E4x4, E1x1/E3x3, E1x1/E4x4, E3x3/E4x4 with all variables to find the one give the best resolution
apply different cut, TOT, TOA, event matching to see if resolution improved

## Slide 3: Target run
Energy resolution of:
1 X 1,
3 X 3,
crystal clustering
will be studied
Run 421~434,481,482 @ HV=45V, C=0xA
Positron beam shoot at Crystal 7
Event selection need to be done with tracking from BM 3&4.
However, tracking requires event matching between BM and PbWO4 beforehand, which I will learn later.
My first try will not include tracking & event selection.
Crystal 7
Channel Mapping

## Slide 4: Crystal status overview
abnormal channels appear in same channel across different run
All channels have basically the same waveform pattern in the same event.
All channels in crystal#3,4,7,10,11 function well.
Problems:
dead channels are found in (exclude in later analysis) crystal#2 : 63, 40, 52, 42 (4 dead) crystal#6: 64, 45, 47, 49, 51 (5 dead) crystal#8: 173, 175 (2 dead) crystal#12: 284, 286, 231, 250 (4 dead)
Abnormal waveform pattern found in crystal#2,6,8,12
Those channels are simply excluded in the analysis
3 4
7
10 11
Well-function crystal

## Slide 5: Run481 vs. Run482: crystal#6
Run#481 Crystal#6
Run#482 Crystal#6
Ch#51
Ch#49
Ch#47
Ch#45
Ch#63
Ch#66 Ch#68 Ch#70
Ch#51
Ch#49
Ch#47
Ch#45
Ch#63
Ch#66 Ch#68 Ch#70
crystal#6: 64, 45, 47, 49, 51 (5 dead)
Upper half of the channel have no problem.
abnormal channels appear in same channel across different run

## Slide 6: Run481 vs. Run482: crystal#8
Run#481 Crystal#8
Run#482 Crystal#8
Ch#175
Ch#173
Ch#175
Ch#173
crystal#8: 173, 175 (2 dead)
Other channels have no problem.
abnormal channels appear in same channel across different run

## Slide 7: ADC saturation issue
Run 422
Run
425
Run
481
Run
423
Beam energy: 823 MeV
Beam energy: 705 MeV
Beam energy: 492 MeV
Beam energy: 584 MeV
Many ADC saturation observed: >700MeV
Those ADC saturation event are included in the following analysis of energy resolution.

## Slide 8: Fitting functions
Gaussian function:
Description: Normal distribution
DSCB
Double-Sided Crystal Ball function:
Description: Double-Sided Crystal Ball consists of a gaussian core that smoothly transitions into a power-law tail on the both sides.
DSCB
DSCB
Gaussian
Resolution definition
Histograms are filled as follow:
take the average of adcMax of all channels within one crystal
Sum of adcMax from step1 over all 3x3 crystals
fill
Say crystal#X has n dead channels:
Then average of adcMAX of all channels is sum of adcMAX over normal channels divided by 16-n
Energy resolution is defined as the ratio of the sigma of the gaussian core to the mean of the gaussian core.
Resolution:
Pedestal peak should be fit as Gaussian by prof Guo

## Slide 9: Run#421 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 8.8%0.2%
Energy resolution: 16%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 8.8%0.2%
Beam energy: 823MeV
ADC
ADC

## Slide 10: Second peak?
This is one event
If sum over 3x3 crystals adcMax average > 2500
Second peak happens when two crystals fire simultaneously?

## Slide 11: Run#422 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
There are only 1929 events in Run#422
Energy resolution: 16%0.1%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 8.9%0.6%
Energy resolution: 8.9%0.6%
Beam energy: 823MeV
ADC
ADC

## Slide 12: Run#481 energy resolution adcMax
3 * 3 energy clustering around crystal#7 give the best energy resolution of 7.7%0.1%
Energy resolution: 7.7%0.1%
Energy resolution: 12%
Energy resolution: 15%
1*1 Crystal#7 (beam center)
Crystal#6,7,8
3*3 Crystal
Beam energy: 705MeV
ADC
ADC
ADC

## Slide 13: Run#482 energy resolution adcMax
1*1 Crystal#7 (beam center)
Crystal#6,7,8
3*3 Crystal
Energy resolution: 8.6%0.1%
Energy resolution: 14%
Energy resolution: 15%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 8.6%0.1%
Beam energy: 705MeV
ADC
ADC
ADC

## Slide 14: Run#423 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 6.4%0.1%
Energy resolution: 14%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 6.4%0.1%
Beam energy: 584MeV
ADC
ADC

## Slide 15: Run#424 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 6.3%
Energy resolution: 14%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 6.3%
Beam energy: 584MeV
ADC
ADC

## Slide 16: Run#425 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 5.8%0.1%
Energy resolution: 013%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 6.3%
Beam energy: 492MeV
ADC
ADC

## Slide 17: Run#426 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 6.0%0.1%
Energy resolution: 13%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 6.0%0.1%
Beam energy: 429MeV
ADC
ADC

## Slide 18: Run#427 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 4.9%
Energy resolution: 12%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 6.0%0.1%
Beam energy: 395MeV
ADC
ADC

## Slide 19: Run#428 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 4.8%0.1%
Energy resolution: 0.13%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 4.8%0.1%
Beam energy: 395MeV
ADC
ADC

## Slide 20: Run#429 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 4.2%0.1%
Energy resolution: 11% 114%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 4.2%0.1%
Beam energy: 297MeV
ADC
ADC

## Slide 21: Run#430 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 4.2%0.0%
Energy resolution: 12%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 4.2%0.0%
Beam energy: 297MeV
ADC
ADC

## Slide 22: Run#431 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 3.1%0.1%
Energy resolution: 9%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 3.1%0.1%
Beam energy: 197MeV
ADC
ADC

## Slide 23: Run#432 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 2.5%0.0%
Energy resolution: 9%0%
3 * 3 energy clustering around crystal#7 give the best energy resolution of 2.5%0.0%
Beam energy: 197MeV
ADC
ADC

## Slide 24: Run#433 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 1.7%0.1%
bad fit
Energy resolution: 0.07 bad fit
3 * 3 energy clustering around crystal#7 give the best energy resolution of 1.7%0.1%
Beam energy: 100MeV
ADC
ADC

## Slide 25: Run#434 energy resolution adcMax
1*1 Crystal#7 (beam center)
3*3 Crystal
Energy resolution: 1.3% 0.0%
bad fit
Energy resolution: 7% 0% bad fit
3 * 3 energy clustering around crystal#7 give the best energy resolution of 1.3% 0.0%
Beam energy: 100MeV
ADC
ADC

## Slide 26: Sanity check
DSCB main peak mean & sigma are linear under beamE scan.
maxPedestal peak mean is not flat (how about the real pedestal?)
DSCB main peak
maxPedestal peak
Second gaussian peak
3x3 maxPedestal

## Slide 27: Energy resolution vs. beam energy
Wrong resolution
Pedestal give low beamE run a relatively large denominator in resolution calculation
Say pedestal = 1000
Resolution = sigma/(mean + 1000)
Wrong resolution
In the graph, DSCB main peak mean is subtracted with maxPedestal mean to calculate resolution.
However, maxPedestal cannot represent the true pedestal.
Resolution =
Resolution =
Solution:
Fit the original ADC waveform and define the pedestal then
Simply subtract the first ADC point while extracting ADCMax