## Slide 1: RPC development weekly progress
Yi-Min Lee, Meng-Hsin Yeh, Yen-Chen Wu
Dept. Physics, National Central University
2026/08/24

## Slide 2: Outline
Cosmic ray trigger setting
  Thin scintillator bar (for individual channel scan)
  Square scintillator bar (cover 3 channels)
High Voltage module usage
Experiment Setup
RPC signal
RPC efficiency test /HV scan

## Slide 3: HV module (CAEN NDT1470 )
High-voltage power supply
Polarization setting
  The output polarity is independently selectable for each channel.
  Lift the side cover, and plug the diode bridge box with the desired symbol (+ or -) facing the connector side.
Max output voltage: 8000V Max output current: 3 mA

## Slide 4: NIM System Module
Signal 1
Signal 2
Model 740 Quad Linear Fan In/Out
Replicates the input signal and distributes it to multiple outputs.
Supports +/ -voltage inputs and provides inverting and non-inverting outputs.
Model 623 Discriminator
Outputs a square wave when the signal exceeds a set voltage threshold. (Threshold 30 mV; Width 60ns)
offset =0
offset =0
Signal 1
Signal 2

## Slide 5: NIM System Module
Signal 1
Signal 2
Phillips Scientific Model 755 (Quad Four-Fold Logic Unit)
The setting of the coincidence level switch then determines whether a logic OR, AND, or Majority logic function will produce an output.
Signal 1
Signal 2
AND

## Slide 6: Scintillator cosmic ray trigger
Setup:
32mm
8mm
162mm
30V
30V
50mm
scintillator | Thin bar | Square bar
Obtain signals | 500 | 1000
Average time per signal | 195.611 s | 10.052s
Standard deviation | 186.835 s | 410.618s
(theoretical by家城)
50mm
Thin bar
8mm
RPC
Square bar
Scintillator1
Scintillator2

## Slide 7: Setup
Scin_top
Scin_bottom
RPC
LV +30V
LV +30V
Amp
LV +5V
HV = +7kV
+HV = -7kV
AND
oscilloscope
signal
RPC 2
RPC 3
RPC 1
Cosmic ray trigger

## Slide 8: Problems
HV & Ground Connection Problem:
Sparks occurred when applying high voltage.
Cause and solution
Electrical tape
Kapton tape
HV Trip / OvC Warning Problem:
The HV tripped and an OvC warning appeared.
Cause and solution
The Iset was set too low (0.3 mA).
Increased ISet from 1 mA to 2 mA.
HV
Ground

## Slide 9: HV Scan
12kV
13kV
14kV
15kV
RPC ch9
RPC ch10
RPC ch8
Cosmic ray trigger
RPC ch9
RPC ch10
RPC ch8
Cosmic ray trigger
RPC ch9
RPC ch10
RPC ch8
Cosmic ray trigger
RPC ch9
RPC ch10
RPC ch8
Cosmic ray trigger

## Slide 10: HV Scan
A lot of noise
16kV
RPC ch9
RPC ch10
RPC ch8
Cosmic ray trigger
50mV
20ns

## Slide 11: Voltage distribution
Unable to find the common threshold of different high voltage.
Planning to observe different threshold of different voltage.
15kV
12kV

## Slide 12: Efficiency
There were a lot of noises when the RPC working
Thershold :
If the volotage over , it won’t be cosidered as signal.
Efficiency is bad.
Manual check
Threshold

## Slide 13: Abnormal channel
RPC ch8
RPC ch9
RPC ch10
Cosmic ray trigger
After HV scan channel 10 has abnormal cutoff
We change its Amp channel ->nothing improve
We change to channel 7 on RPC, signal becomes normal.
RPC channal 10 seems to have some problem

## Slide 14: Summary
Finish the bar and square scintillator trigger test
Apply HV on RPC (12,13,14,15,16kV)