## Slide 1: Particle Detectors for Medical Applications
Chih-hsun Lin
Institute of Physics, Academia Sinica

## Slide 2
Radiation
is
transmission of energy through space or materials.
Electromagnetic waves
Charge or neutral particles
*https://en.wikipedia.org/wiki/Radiation

## Slide 3
Particle Physics is the study
of
fundamental particles and forces.
Detectors developed in particle experiments could be useful in radiation applications.
*https://en.wikipedia.org/wiki/Particle_physics

## Slide 4: Discovery of Radiation
In 1895, Wilhelm Röntgen discovered X-ray by experiments with Crookes tubes.
In 1896, Henri Becquerel discovered that uranium salts emitted rays with penetrating power.
In 1898, Pierre and Marie Curie discovered Polonium (210Po) and Radium (226Ra).
The term “Radioactivity” was coined by Marie Curie。
Notes: H->ZZ->eeuu

## Slide 5: Application of Radiation
First X-ray image of Wilhelm Röntgen's wife's hand.
Great tool in the medical examination.
Notes: H->ZZ->eeuu

## Slide 6
Taking an X-ray image with early Crookes tube

## Slide 7: Radiotherapy – Early Stage
X-ray
  1896- treated a local relapse of breast carcinoma (Emil Grubbe in Chicago)
  1896 - treated stomach cancer (Victor Despeignes in Lyon, France)
  1896- treated a skin tumor (Léopold Freund in Vienna, Austria)
Radium
  Skin, prostate, and gynecologic cancers.
Notes: H->ZZ->eeuu

## Slide 8: Wrong Application – X-ray Fitting
Notes: 1970才正式禁止

## Slide 10: Radiosensitivity
Radiation affects both normal cells and cancel cells.
  Different types of cells have different radiosensitivity. They are more sensitive if
    Cells are highly proliferative with rapid cell divisions.
    Less differentiated like stem cells
  Cancel cell is usually more sensitive to radiation.
In general, radiotherapy must
  Deliver less radiation to normal tissues.
  Deliver sufficient radiation to cancel cells.
Precision radiation delivery and verification are important.
Good Radiation Detector is required.

## Slide 11: Radiation Delivery
Body surface:
  Electrons from Linac
Inside body:
  X-ray from Co60 or Linac
  Protons from cyclotron or Synchrotron
Dose depth in water

## Slide 12
Bragg curve, discovered by Sir William Henry Bragg in 1903.
Particle therapy was proposed by Robert R. Wilson in 1946.
Notes: Ra226 decay mode Ra226-> Alpha + Rn222
Bragg 1915 Nobel Prize in Physics: "for their services in the analysis of crystal structure by means of X-rays".

## Slide 13
Proton
X-ray
Notes: Tricho sales corporation established in 1924

## Slide 14: Cyclotron Machine
Harvard Cyclotron Lab: 165 MeV (1956), ~ 1000 tons
Notes: 治療到2002

## Slide 15: Particle Therapy Center in US
1990, First hospital-based facility at
the Loma Linda University Medical Center in California

## Slide 16: IBA cyclotron for proton therapy
There are more than 100 particle therapy centers in the world.
More than 30,000 patients treated per year.

## Slide 17: Proton Beam Delivery – Passive
Broad Proton Beam  Dose distribution in various conditions
Ionization detector

## Slide 18: Gaseous Ionization Detector
*https://www.radiation-dosimetry.org/what-is-proportional-region-ionization-detector-definition/
R.T.
Proportional
chamber
Geiger counter
Spark
chamber
Many particles or X-rays
  Bigger signal
  Stability
Ionization charges depend on
  Chamber volume
  Particle Energy
  Particle type
Precision dose measurement

## Slide 19: Gaseous Ionization Detector
*AEC Radiation Measuring Instruments
*Physics in Medicine and Biology 60(22):8625-8641

## Slide 20: 2D Gaseous Ionization Detector
200 MeV Proton Beam and Collimator Diameter = 10 cm

## Slide 21: 100 MeV Proton Beam in Water
100 MeV Proton Beam in water

## Slide 22: Pencil Beam Scanning
Active Scanning

## Slide 23: Pencil Beam Scanning

## Slide 24: Characteristics of Pencil Beam Scanning
High precision
  Small beam size ( ~ few mm)
  Fast scanning (Max speed = 20 m/sec)
High flexibility
  Use full beam energy = Maximum treatment depth.
  Adjustable scanning speed and beam current
Bring new requirements to the radiation detector

## Slide 25: Quality Assurance Requirements
Position
Profile
Intensity
Scanning speed
Energy
Large area
One detector
Short QA time
1.88 mm
0.12 mm
1
2
3
4
5
1 : HV board
2 : Spacer
3 : Strip detector
4 : Spacer
5 : HV board
Fine Strip X&Y Strip ion detector
Fine pitch but large area.
Less channel numbers to reach high rates in measurement.
Wedge phantom for energy measurment

## Slide 26: CROSS Detector
Type: Strip of Parallel Plate Ionization Chambers
Active area: 330 x 330 mm2.
Maximum sampling rate: 20 kHz One measurement per 50 sec
Single interface: USB 2.0

## Slide 27: Beam position and spot size measurement (single event)
Peak position = -74.0 mm
s = 4.64 mm
X
Y
Peak position = 73.1 mm
s = 4.52 mm
Beam Energy = 150MeV

## Slide 28: Position measurement uncertainty
 x~ 0.09 mm
Histogram of position measurement
Position vs time(ms)

## Slide 29: Proton Beam Trajectory (30x30 cm)
Scanning Plan
Measurement

## Slide 30: Proton Beam Trajectory (30x30 cm)

## Slide 31: Scanning Speed
Chart: Time (msec)

## Slide 32: Scanning Speed
Chart: Time (msec)
14 | -19.401 | -17.793 | -17.455 | -15.811 | -15.459 | -13.868 | -13.515 | -11.923 | -11.627 | -9.904 | -9.705 | -7.948 | -7.582 | -5.994 | -5.702 | -4.138 | -3.75 | -2.227 | -1.704 | 0.144 | 0.779 | 2.206 | 2.476 | 4.191 | 4.608 | 6.072 | 6.409 | 7.919 | 8.326 | 9.871 | 10.287 | 11.819 | 12.304 | 13.792 | 14.292 | 15.817 | 16.289 | 17.783 | 18.224 | 19.834 | 20.257
V = 19.83±0.09 mm/msec

## Slide 33: Beam Energy Verify (with Wedge)
Chart: 
| -80 | 80
-80 80 | 0 | 0
PBS Scan direction
Thin
Thick

## Slide 34: Beam Energy Verify ( 70MeV – 220MeV )
Chart: Normalize_Energy
| 22.45 | 22.98 | 23.49 | 23.98 | 24.46 | 24.97 | 25.42 | 25.84 | 26.26 | 26.67 | 27.05 | 27.39 | 27.71 | 27.99 | 28.22 | 28.41 | 28.59 | 28.72 | 28.84 | 28.94 | 29.04 | 29.12 | 29.18 | 29.25 | 29.3 | 29.33 | 29.38 | 29.42 | 29.44 | 29.46 | 29.48 | 29.51 | 29.52 | 29.53 | 29.54 | 29.35 | 29.33 | 29.51
70MeV | 78.65 | 80.5 | 82.79 | 84.86 | 87.29 | 89.99 | 91.73 | 93.75 | 95.42 | 97.44 | 98.6 | 99.09 | 100 | 99.57 | 98.81 | 96.76 | 93.31 | 90.76 | 86.26 | 81.35 | 76.03 | 69.88 | 64.56 | 58.35 | 52.11 | 46.46 | 40.39 | 34.9 | 30.15 | 25.26 | 21.72 | 17.8 | 15.05 | 12.44 | 10.43 | 8.51 | 6.9 | 5.88
80MeV | 52.95 | 53.4 | 54.31 | 54.82 | 55.53 | 56.65 | 56.92 | 57.96 | 58.79 | 59.37 | 60.43 | 61.5 | 62.43 | 63.91 | 64.58 | 66.18 | 67.47 | 68.94 | 70.54 | 71.59 | 73.26 | 75.02 | 77.96 | 79.84 | 81.82 | 84.66 | 87.01 | 89.7 | 92.22 | 94.08 | 96.4 | 97.95 | 99.51 | 100 | 99.6 | 98.73 | 97.49 | 94.17 | 89.98 | 85.35 | 80.34 | 73.09 | 67.06 | 60.74 | 53.14 | 46.16 | 40.25 | 34.6 | 28.79 | 23.45 | 19.78 | 16.26 | 12.99 | 10.52 | 8.53 | 7 | 5.61 | 4.51

## Slide 37: Summary
Ionization detector is an important QA tool of radiotherapy.
Scintillating detector is also a tool by taking an advantage of CMOS imaging sensor.
Notes: H->ZZ->eeuu